# Risk analysis

# Noora Alhajeri

Product title: Wearable IOT Based device for critical care Alzheimer’s patients

| energy | chemical and biological | operational | information |
| --- | --- | --- | --- |
| uncovered wires | material and fabric used may cause rashes and allergic responses | improper placement of sensors | unproper size of the hand band |
| electric shock while recharging |  | wrong connections | inaccurate measures and recordings |
| burns due to overheated battery |  | incorrect codes for the sensors | improper positioning of sensors |
|  |  | damaged circuit |  |
|  |  | poor detection of location |  |

Table1: Project hazard identification

| Hazard | Hazard situation | Harm |
| --- | --- | --- |
| uncovered wires | wires touch the skin | electrical injury |
| electric shock while recharging | electrical shock due to excessive leakage currents | electrical injury |
| burns due to overheated battery | skin with contact with battery | Burns and pain |
| material and fabric used may cause rashes and allergic responses | allergic skin | irritated skin and itchy rash |
| wrong connections | improper functionality | wrong measurements |
| poor detection of location | lost patient and panic of caregivers | possible accidents may occur to the patient |
| damaged circuit | sensors will not work | wrong measurements or no measurements |
| incorrect codes for the sensors | sensors will not work | wrong measurements or no measurements |
| improper positioning of sensors | sensors not able to detect signs | wrong measurements or no measurements |
| improper size of the hand band | sensors not able to detect signs | wrong measurements or no measurements |
| device does not work properly | improper detection of vital signs | wrong measurements or no measurements |

| Common terms | Possible description |
| --- | --- |
| Catastrophic | the hazard can cause death and illness. |
| Critical | the hazard can result in long-term disability or a life-threatening injury. |
| Moderate | the hazard can result in serious injury or illness, property, or equipment damage. |
| Marginal | the hazard could cause illness, injury, or equipment damage but its effects would not be serious. |
| Negligible | the hazard will not result in serious injury or illness or has a remote possibility of damage. |

Table2: Project hazardous situations

Table3: five quantitative severity levels

| Common terms | Possible description |
| --- | --- |
| frequent | 1 to 1000 |
| Reasonably probable | 1 in 100000 |
| Remote | 1 in 10000000 |
| Extremely remote | 1 in 1000000000 |
| Extremely improbable | 1 in 10000000000 |

Table4: semi-quantitative probability levels

| Risk # | Hazard | Probability of occurrence | Hazard situation | Harm | Severity of harm |
| --- | --- | --- | --- | --- | --- |
| R1 | uncovered wires | Reasonably probable | wires touch the skin | electrical injury | Critical |
| R2 | electric shock while recharging | Reasonably probable | electrical shock due to excessive leakage currents | electrical injury | Critical |
| R3 | burns due to overheated battery | Extremely remote | skin with contact with battery | Burns and pain | Critical |
| R4 | material and fabric used may cause rashes and allergic responses | Remote | allergic skin | irritated skin and itchy rash | Catastrophic |
| R5 | wrong connections | Frequent | improper functionality | wrong measurements | Marginal |
| R6 | poor detection of location | Extremely remote | lost patient and panic of caregivers | possible accidents may occur to the patient | Marginal |
| R7 | damaged circuit | Reasonably probable | sensors will not work | wrong measurements or no measurements | Moderate |
| R8 | incorrect codes for the sensors | Remote | sensors will not work | wrong measurements or no measurements | Marginal |
| R9 | improper positioning of sensors | Frequent | sensors not able to detect signs | wrong measurements or no measurements | Marginal |
| R10 | improper size of the hand band | Extremely improbable | sensors not able to detect signs | wrong measurements or no measurements | Negligible |
| R11 | device does not work properly | Reasonably probable | improper detection of vital signs | wrong measurements or no measurements | Marginal |

Table5: Estimated risk for the project

|  | Negligible | Marginal | Moderate | Critical | Catastrophic |
| --- | --- | --- | --- | --- | --- |
| frequent |  | R5, R9 |  |  |  |
| Reasonably probable |  | R11 | R7 | R1, R2 |  |
| Remote |  | R8 |  |  | R4 |
| Extremely remote |  | R6 |  | R3 |  |
| Extremely improbable | R10 |  |  |  |  |

Table6: semi-quantitative 5X5 risk matrix

|  | Negligible | Marginal | Moderate | Critical | Catastrophic |
| --- | --- | --- | --- | --- | --- |
| frequent |  | R5, R9 |  |  |  |
| Reasonably probable |  | R11 | R7 | R1, R2 |  |
| Remote |  | R8 |  |  | R4 |
| Extremely remote |  | R6 |  | R3 |  |
| Extremely improbable | R10 |  |  |  |  |

Table7: semi-quantitative 5X5 evaluation matrix

| Risk # | Hazard | Hazard situation | Acceptable risk? | Protective control measures | Acceptable RR? |
| --- | --- | --- | --- | --- | --- |
| R1 | uncovered wires | wires touch the skin | NO | Utilize insulated wires | YES |
| R2 | electric shock while recharging | electrical shock due to excessive leakage currents | NO | Utilize appropriate case to isolate the battery | YES |
| R3 | burns due to overheated battery | skin with contact with battery | NO | Reduce pressure on battery | YES |
| R4 | material and fabric used may cause rashes and allergic responses | allergic skin | NO | State in the user manual the material used in the hand band with restating the presence of allergenics | YES |
| R5 | wrong connections | improper functionality | YES | NA | NA |
| R6 | poor detection of location | lost patient and panic of caregivers | YES | NA | NA |
| R7 | damaged circuit | sensors will not work | NO | Utilize insulated material | YES |
| R8 | incorrect codes for the sensors | sensors will not work | YES | NA | NA |
| R9 | improper positioning of sensors | sensors not able to detect signs | YES | NA | NA |
| R10 | improper size of the hand band | sensors not able to detect signs | YES | NA | NA |
| R11 | device does not work properly | Table8: Risk control measures for the project use improper detection of vital signs | YES | NA | NA |